## Slide 1: Sensor Registration: Procurement and Build
The board, its adapter and the fixtures needed for the registration captures
Procurement and mechanical build · October 2026
Notes: This is the first of two decks on the registration captures, the session that gathers the sample set for registering the 3D sensor to the robot from plane-to-plane correspondences. It covers what has to exist before a capture session: the flat board, its adapter on the robot flange, the run-out fixture and the measuring tools, with estimated costs and candidate suppliers. Almost all of it is the stage-1 calibration hardware reused; the deck says what to check rather than what to make, and carries the specifications so that the items can be made if stage 1 has not been run. The second deck covers the test procedure itself and what the robot program must do. Both summarize the registration capture procedure in docs/procedures of the plane_plane_registration repository, whose section 1 and appendices A and D hold the full details.

## Slide 2: One flat board carries the known geometry
The board's front face as the sensor sees it: the tool frame at the center, the four places where D is measured, and the three-point mounting behind it.
Board
200 x 150 mm plate, at least 6 mm thick, matte light gray, flat to 0.05 mm after painting and mounting
Board adapter
Doweled plate on the ISO flange; three support pads, three edge pins, three clamp fingers hold the board square to the flange axis
Tool frame
Origin at the center of the front face, z along the board normal toward the sensor, x along the long edge; the robot reports this frame
Notes: Registration needs only one target: a flat board whose plane the sensor can measure and whose pose the robot reports. The sensor fits a plane to the board in each capture; the robot's reported tool frame gives the same plane in the robot's base frame; the transform that maps one set of planes onto the other is the registration. No spheres and no three-ball nest are needed, which is why this deck is short. The board, its adapter and its tool frame are the stage-1 ones (procedure section 1a and section 3, figure 1). The front face is what matters: its flatness, its matte finish and the perpendicularity of its mounting on the flange axis.

## Slide 3: What must be built
Item | Quantity | Drawing | From stage 1
Flat target (board) | 1 | SC1-05 (outline) | yes
Board adapter: flange plate with pads, edge pins and clamp fingers | 1 | SC1-05 | yes
Run-out fixture: dial indicator on a magnetic base and steel base plate | 1 | none needed | yes
Nothing new if stage 1 was run. Section 1b of the procedure; the rigid sensor mount is in hand from the existing drawing (1a). Each item carries its specification and drawing so it can be made for this procedure alone.
Notes: This is the complete list of apparatus the session needs beyond the sensor, the robot and the computers, as section 1b of the procedure gives it. Every item was built for the stage-1 calibration session and is reused as it is: the same board, the same adapter, the same run-out fixture, the same sensor mount. If stage 1 has not been run, the items are built to the stage-1 specification and drawing, which the procedure carries: the plate specification in section 1d, the run-out fixture in 1e, and drawing SC1-05 for the board adapter in appendix D, copied from the stage-1 repository. The rigid sensor mount is built from the existing drawing from previous work and is not drawn again.

## Slide 4: What must be bought
Item | Purpose | Estimated cost (USD) | From stage 1
Calipers with depth rod, 150 mm | Board distance D; board size | $30 to $150 | yes
Straightedge 300 mm and feeler gauges | Board flatness after painting and mounting | $75 to $240 | yes
Thermometer | Room temperature in the session notes | $15 to $40 | yes
Machinist's square | Board x axis against the adapter | $30 to $80 | yes
Consumables | Isopropyl alcohol and wipes, shim stock | $20 to $60 | no
A phone camera is enough for the setup photos. Already in hand: sensor, robot, capture computer with its communication software, analysis computer.
Notes: These are the measuring tools and consumables of section 1c of the procedure. All of the tools were bought for stage 1; only consumables are new. The calipers measure the flange-to-face distance D, which sets the tool frame and goes in the session notes, and the board half-sizes, which go in the pose log; the straightedge and a 0.05 mm feeler leaf check the board's flatness after painting and after mounting; the square sets the board's x axis when the tool frame is defined; the thermometer gives the room temperature for the session notes. No micrometers are needed, since there are no spheres to measure. The sensor, the robot, the capture computer with its communication software and the analysis computer are already in hand (section 1a).

## Slide 5: New spending: $20 to $60 after stage 1
Item to build | Estimated cost (USD)
Flat target (board) | $50 to $400
Board adapter | $250 to $600
Run-out fixture: dial indicator, magnetic base, steel base plate | $240 to $570
Rigid sensor mount | not estimated
$20 to $60
new spending, everything else in hand from stage 1
$710 to $2,140
built from nothing: $540 to $1,570 to build, $170 to $570 to buy
Estimates of October 2026 from the stage-1 procedure, plus or minus 50 percent; the board adapter is the least certain figure.
Notes: The cost figures are the stage-1 procedure's, kept in one script (docs/procedures/build/costs.py) from which the procedure's section 1c and this slide are both filled, so they cannot disagree. They come from suppliers' list prices where these exist and from typical United States job-shop rates of about 80 to 150 dollars an hour otherwise. If the stage-1 session has been run, the board, the adapter, the run-out fixture and the tools exist, and the only new spending is consumables. Built from nothing, the three items come to 540 to 1,570 dollars and the tools to 170 to 570 dollars. Treat every figure as plus or minus 50 percent and get quotes; the board adapter, a one-off machined plate with lapped pads, is the least certain.

## Slide 6: Purchase specification: flat plate
Requirement | Flat plate (board)
Size | 200 x 150 mm, at least 6 mm thick
Material | Finish-ground aluminum tooling plate (for example MIC-6) or float glass
Flatness | 0.05 mm over the front face, checked with a straightedge and a 0.05 mm feeler leaf after painting and after mounting
Front face | Thin, even matte light-gray paint; do not bead-blast a plate thinner than about 10 mm
Edges | Bottom long edge and left short edge square and clean: they rest on the adapter's edge pins
Documents | Supplier's flatness report, kept with the session notes
Locating edges
Support pads
Procedure section 1d; the board outline is on drawing SC1-05.
Notes: This is the purchase specification of section 1d of the procedure, identical to the stage-1 plate, to hand to a supplier or a grinding shop. The board is checked flat after painting and after mounting because that is the state the sensor sees. Thickness is not a stiffness matter: a 6 mm plate sags only about 0.002 mm under its own weight; the plate only has to survive the finish and the mounting flat. Bead-blasting is ruled out below about 10 mm because peening one face bows the plate; a thin, even matte paint over the ground or glass surface is the finish. The two reference edges rest against the adapter's edge pins, so they must be square and clean.

## Slide 7: Candidate suppliers
Item | Candidate suppliers | Notes
Board plate | McMaster-Carr MIC-6 plus a grinding shop; Edmund Optics float glass | Ground or glass to 0.05 mm with a flatness report; painted matte; re-checked in-house after painting and mounting
Dial indicator and magnetic base | Mitutoyo 2046 series; Noga MG71003 or DG-61003 | For the run-out fixture; McMaster-Carr and the usual tool suppliers
Calipers, straightedge, feeler gauges | Mitutoyo or Starrett; DIN 874 grade 0 or better straightedge | Transcat, MSI-Viking, McMaster-Carr
Board adapter and run-out base plate | Local machine shop | Drawing SC1-05; the base plate is a plain drilled plate
From suppliers' web pages, October 2026; not an endorsement. Confirm flatness, finish and price before ordering.
Notes: This is appendix A of the procedure in brief, the stage-1 supplier list reduced to the items this procedure uses. The plate is one of the two materials of the specification on the previous slide: MIC-6 tooling plate is sold with a stated flatness of about 0.13 mm, coarser than the 0.05 mm asked for, so it is ordered oversize and finish-ground by a local shop; float glass is flat enough as delivered and is the better choice where no grinding shop is at hand. In both cases the supplier's or grinding shop's flatness report is the starting document, and the 0.05 mm is confirmed in-house with the straightedge and feeler leaf after painting and after mounting, the state the sensor sees. The instruments are stock items. The adapter is a one-off job for a machine shop, built to drawing SC1-05.

## Slide 8: Run-out fixture: indicator and mounting
The run-out check: joint 6 turns the flange, the indicator stays still, the reading must stay within 0.05 mm.
Dial indicator
0.01 mm graduation, about 10 mm travel; for example Mitutoyo 2046 series, about $50 to $150
Magnetic base
Articulating, with fine adjustment; for example Noga MG71003 or DG-61003, about $140 to $320
Steel base plate
About 150 x 100 x 12 mm, bolted to the cell table with two M8 screws, so the base holds on a non-magnetic table and does not creep
Placement
Where the indicator tip reaches the board's front face about 20 mm from an edge with the robot in its run-out pose
Notes: The run-out check in deck 2 needs the dial indicator held still while the robot turns the board about the flange axis; it tests whether the board face is perpendicular to that axis, which the tool frame's z axis assumes. A magnetic base alone can creep on a non-magnetic table, so the fixture stands on a steel plate bolted down. The stage-1 fixture serves as it is; section 1e of the procedure gives the parts and their placement. The figure is figure 1a of the procedure: the flange, the adapter and the board in section, with the indicator tip on the front face.

## Slide 9: Keep the board flat after mounting
The front face: tool frame axes, the path of the indicator tip, the four D places, and the adapter's pads, pins and clamp behind the board.
Plate
200 x 150 mm, at least 6 mm thick; ground aluminum tooling plate or float glass
Finish
Matte light gray paint on the front face; do not bead-blast a plate thinner than about 10 mm, it bows
Flatness
0.05 mm, checked after painting and mounting, not before; keep the flatness report
Adapter
Three support pads, three edge pins and three clamp fingers hold the board square to the flange axis, centered on it, and in the same place after a remount
Notes: The adapter is the one part that needs a drawing. Its three lapped pads define the board plane, its three edge pins locate the board within that plane, and its three clamp fingers hold the board against the pads without bending it, so that a remount puts the board back in the same place. The plane of the pads must be parallel to the adapter's back face, which sits on the robot flange; that is what the run-out check verifies after mounting. Drawing SC1-05 in appendix D of the procedure gives every dimension and tolerance; the drawing is generated in the stage-1 repository and copied here. The figure is figure 1b of the procedure.

## Slide 10: Checks before the board goes to the robot cell
Board adapter made to drawing SC1-05, flange interface confirmed against the robot's flange drawing
Board flat to 0.05 mm after painting and after mounting on the adapter; flatness report kept
Front face matte, uniform and free of highlights; reference edges square and clean
Board and adapter weighed for the tool load data
Run-out fixture's base plate bolted to the cell table where the indicator reaches the board
Dowels seated; after any remount, the remount check of the procedure (run-out, D, reference capture) passes
Notes: These are the build and preparation checks of sections 1 and 3 of the procedure, gathered in one place. Each protects something the software cannot repair later: a board that is not flat gives plane residuals the registration reads as sensor error; a glossy spot returns a highlight and a bad read; a loose or unseated adapter shifts the board between poses and spoils every pose after it; wrong tool load data shifts every reported pose. The run-out and D measurements themselves are session steps, in deck 2. A remount is verified, not assumed: section 3 of the procedure repeats the run-out check and the D measurement and re-captures a saved reference pose, so that any change of the board's plane relative to the flange shows before it spoils a run.